| Depth cm | Field Size cm2 | Highest Relative difference % Between MCNP and Measured in the Flat Region* |
| --- | --- | --- |
| 1.5 (d max) | 5 x 5 | 1.13 |
| 5 | 5 x 5 | 2.05 |
| 10 | 5 x 5 | 1.01 |

The flat region is defined by subtracting 1 cm from both sides of the field.

| Depth cm | PDD% (for 5 x 5 cm2 Field Size at 100 cm SSD) | Highest Relative difference % Between MCNP and Measured in the Decay Region* |
| --- | --- | --- |
|  | MCNP | Measured |  |
| 5 | 86.58 ± 0.61 | 85 | 1.85 |
| 10 | 65.6 ± 0.48 | 63.8 | 2.77 |

The Decay Region: is defined as the region after dmax (1.5 cm for 6 MV)