- Accuracy test:

| requirements | upon installation and annually |
| --- | --- |
| limits | 10% |
| Test passed/failed | passed |

| source | Cs-137 (30y) |
| --- | --- |
| Measured activity #1 | 175.3 uCi |
| Measured activity #2 | 174.1 uCi |
| Measured activity #3 | 174.9 uCi |
| Average | 174.7 uCi |
| Predicted | 179.3 uCi |
| Accuracy (limit ±10%) | 2.5 % |
| Test Passed/failed | Passed |

| source | Co-57 (0.74y) |
| --- | --- |
| Measured activity #1 | 188.6 uCi |
| Measured activity #2 | 187.6 uCi |
| Measured activity #3 | 188.6 uCi |
| Average | 188 uCi |
| Predicted | 192 uCi |
| Accuracy (limit ±10%) | 1.9 % |
| Test Passed/failed | Passed |

| source | Ba-133 (10.74y) |
| --- | --- |
| Measured activity #1 | 211 uCi |
| Measured activity #2 | 210 uCi |
| Measured activity #3 | 210 uCi |
| Average | 210.3 uCi |
| Predicted | 209uCi |
| Accuracy (limit ±10%) | 0.6 % |
| Test Passed/failed | Passed |