Energy resolution (without collimator) (SEMINSE e.cam); );(daily and before each new radionuclide used)

-using the same set as in (peaking).

-At the spectrum curve, measure the FWHM, and then apply in the equation;

* 100

The photo peak should be at 140Kev fro Tc-99m.and the resolution limits are 8%-12%.