3-intrinsic Resolution test (without the collimator) ;( weekly, after repair)

1-matrix size is set to its maximum value 512*512.

2-Tc-99m 500uci is used as a point source at a distance about 5*FOV, a whole count of 10*10^3kc is enough.

3- A resolution phantom as in the fig is placed on the crystal of each head “detector”, in this test we‘ve used only on e head #1, the other was used to measure the linearity.

4-the test pattern should be rotated to get the resolution in other quadrants.

5-the smallest pattern that can be resolved is recorded.

6-accepted limits are 2.5 – 3.5 mm.

# Text boxes

3.5mm

2mm

2.5mm

3mm