1-field uniformity:

- Intrinsic uniformity (without collimator) ;( daily check)

It measures the performance of the crystal.

Procedures:

- A point source is placed at a distance of about 5 *FOV(field of view),150cm.

- Source specification , Tc-99m 148uci

- Using the built I calibration setting in the camera software, in order to make a dome correction.

- A dome correction is made because the point source is facing both camera detectors, so their position is not flat, but tilted as in the fig.

- A glass mask is used to cover the PM tubes at the edge which might affect the reading.

6-total count is set to 30*10^6counts and we shouldn’t exceed 20kc/s.

7-count rate recorded by D1 is 19kc/s, count rate recorded by D2 is 19.16kc/s.

8- Matrix size 256*256, where the recommended is (64*64).

9-the integral and differential uniformity of the UFOV & CFOV are calculated   for both   detectors acquired by available camera software, and it is compared with the specifications.

# Text boxes

Point source

Detector 1

Detector 2

Pb edge

Glass