- constancy test:

Using the exponential decay expression to calculate the predicted value:

| source | Initial activity A0 At (1-5-2003) | Predicted at 19-11-2006 After (3.54 y) | Measured | Percentage difference = |
| --- | --- | --- | --- | --- |
| Ba-133(10.74y) | 263.1uci | 209uci | 211uci | ≈1% |
| Cs-137(30y) | 194.6 uci | 179.3 uci | 176uci | 1.8% |
| Co-57(270.9d or 0.74 y) | 5330uci | 193.6 uci | 189 uci | 2.3% |

| using Cs-137 in all of these channels |
| --- |
| channel | activity uci | predicted uci | percentage difference % |
| Cs-137 | 176.00 | 179.30 | 1.84 |
| co-57 | 296.00 | 299.90 | 1.30 |
| I-131 | 225.00 | 227.86 | 1.26 |
| In-111 | 128.50 | 130.18 | 1.29 |
| Ga-67 | 305.00 | 309.50 | 1.45 |
| Xe-133 | 191.80 | 194.20 | 1.24 |
| I-123 | 130.40 | 132.00 | 1.21 |
| Tl-201 | 190.10 | 193.20 | 1.60 |
| Mo-99 | 1815.00 | 1846.00 | 1.68 |
| Tc-99m | 335.00 | 342.27 | 2.12 |

| requirements | Daily checked |
| --- | --- |
| Limits (10%) | All measurements are within the limits |
| Test passed/failed | passed |