4-Linearity (without the collimator);(weekly, after repair)

1-matrix size is set to its maximum value 512*512.

2-Tc-99m 500uci is used as a point source at a distance about 5*FOV, a whole count of 10*10^3kc is enough.

3- A bar pattern phantom as in the fig is placed on the crystal of each head “detector”, in this test we‘ve used only one head #2, the other was used to measure the intrinsic resolution.

4- Using the software tools, scan tow adjacent lines at the center, bottom, and the top of the field.

5-obtain the profile of each one of the previous scan, and measure the distance between the FWAMs of the tow profiles.

6-compare the result with the actual distance.

# Text boxes

30cm

30mm between adjacent lines